What it is:

An analog X-Y pen plotter output to HP-GL translator. (Hewlett-Packard Graphics Language (http://en.wikipedia.org/wiki/HPGL)

What it does:

Some measurement devices, like the Tektronix 2232 Digital Storage Oscilloscope, have an analog X-Y pen plotter output port so that the display image can be captured to an analog X-Y pen plotter or recorder. See oscilloscope display, below.

The problem is one wishes to capture the display image in digital form onto a modern day personal computer, since analog pen plotters are virtually extinct. PC images can be pasted directly into documents.

The 'Analog X-Y Pen Plotter Output To HP-GL Translator' will continually sample the X and Y analog voltages from the plotter port, digitize them, and package the values into a series of plain text HP-GL plotter commands, which are output via a USB port to a file on a PC. The HP-GL file can be processed into a display image by any commonly available plotter emulator application, such as PrintCapture (http://www.printcapture.com). See actual plot image, below.

What it's made of:

- Arduino UNO for the ADC conversions, digital filtering, assembly of HP-GL commands and serial output.

- Arduino 'Shield' type board for signal conditioning (voltage translation), voltage reference and power monitoring. See schematic and boards, below.

- DB9 cable for connection to the oscilloscope plotter port. See Auxiliary Connector, below.

- USB cable between Arduino UNO and PC.

- 9VDC AC adapter (a wall wart).

What's on the Arduino shield board:

The Arduino shield board contains circuitry comprised of (see schematic);

- Two (2) precision, unity-gain differential amplifiers (wired as summing amps) to translate the ±2.5V analog X and Y voltages into 0 to 5V for the Arduino analog pins. U2, U3

- A 'charge pump' voltage converter to convert 9VDC to ±9VDC to power the amplifiers. U1, C1, C2

- A precision 2.50V voltage reference to the amplifiers for the voltage translation. U4

- A precision 5.00V voltage reference to the Arduino AREF pin. U5

- A voltage divider network so the 9V Vin can be reduced to 5V for the Arduino analog pin. A schottky diode protects the pin if a higher voltage power supply is used. R1, R2, D1

- A bi-color red/green LED as a Go/No Go indicator that the 9V power supply is connected, or not. The Arduino sketch will stay in the setup() routine until a 9VDC, or greater, power supply is plugged into the Arduino UNO. These IC's cannot run accurately on the USB ~5V power! LED, R3

- A slide switch to start the data stream and stop it when the plot is completed. S1

- A pair of capacitors to filter the DAC dithering of the plotter port X and Y voltage signals. C9, C10

- A 10-pin header for connection to the oscilloscope plotter port. J1

Operation:

- Slide switch to OFF.

- Connect Arduino via DB-9 connector to oscilloscope plotter port.

- Connect Arduino USB to PC.  Arduino powers up, LED is red.

- Connect 9VDC power to Arduino.  LED is green.

- Launch listening application, like TeraTerm or PrintCapture. Ensure your application is set to the Arduino UNO COM port 6, 115200, 8, none, 1.

- Slide switch to ON. (data stream commences)

- Immediately press PLOT key on scope.

- When scope plot completes immediately slide switch to OFF. (data stream halts)

- Reset Arduino in order to start the next plot.

Example HP-GL output to file:

IN;

PS8900,8900;

IP0,0,8900,8900;

SC0,250,0,1023,1;

PU;          // Pen Up

SP1;

PA77,984;    // Plot Absolute to (77,984)

...          // more data points (PA) every 1.6 milliseconds

PD;          // Pen Down

PA53,92;

...          // more data points (PA) every 1.6 milliseconds

PU;

SP;

IN;

Actual plot image from an HP-GL file created by the Translator of a Tektronix 2232 oscilloscope display. Image created by PrintCapture.

Actual Tektronix 2232 oscilloscope display.

Schematic:

Arduino shield board front

Arduino shield board back

Tektronix Auxiliary Connector for X-Y plotter. Relay is for pen up/down

Arduino Sketch:

/*

****************************************************************************************|

* Tektronix 2232 Digital Storage Oscilloscope analog X-Y pen plotter to HP-GL

* translator. (Hewlett-Packard Graphics Language http://en.wikipedia.org/wiki/HPGL)

*

* John Horstman

* 3/19/2017 Rev H

*

* This sketch continuously reads the X and Y analog pen plotter voltages, and the

* pen up/down signal, and packages the ADC values in successive HP-GL commands. The

* commands can be serially output to a file, via TeraTerm, or to a plotter application

* like PrintCapture at http://www.printcapture.com/.

*

* The HP-GL commands used in this sketch are:

* IN;  Initialize

* PS:  Plot size

* IP;  Input P1 & P2

* SC;  Scale

* SP;  Select pen

* PU;  Pen UP

* PD;  Pen DOWN

* PA;  Plot Absolute

*

* Example HPGL output of this sketch:

* IN;

* PS8900,8900;

* IP0,0,8900,8900;

* SC0,250,0,1023,1;

* PU;

* SP1;

* PA77,984;

* ...          // data points every 1.6 milliseconds

* PD;

* PA53,92;

* ...

* PU;

* SP;

* IN;

*

* This sketch is for Arduino UNO with Proto Shield. The Proto Shield contains signal

* conditioning circuitry comprised of;

* 1. Two (2) precision, unity-gain differential amplifiers (wired as summing amps)

*    to translate the Â±2.5V analog X and Y voltages into 0 to 5V for the Arduino analog

*    pins.

* 2. A 'charge pump' voltage converter to convert 9VDC to Â±9VDC to power the amplifiers.

* 3. A precision 2.50V voltage reference to the amplifiers for the voltage translation.

* 4. A precision 5.00V voltage reference to the Arduino AREF pin.

* 5. A voltage divider network so the 9V Vin can be reduced to 5V for the Arduino

*    analog pin. A schottky diode protects the pin if a higher voltage power supply

*    is used.

* 6. A bi-color red/green LED as a Go/No Go indicator that the 9V power supply

*    is connected, or not. This sketch will stay in the setup() routine until a

*    9VDC, or greater, power supply is plugged in to the Arduino. These IC's

*    cannot run accurately on the USB ~5V power!

* 7. A slide switch to start the data stream and stop it when the plot is completed.

* 8. A pair of capacitors to filter the DAC dithering of the plotter X-Y voltage signals.

* 9. A 9-pin cable with DB-9 male connector for connection to the oscilloscope plotter

*    port.

*

* Operation:

* Slide switch to OFF.

* Connect Arduino via DB-9 connector to oscilloscope plotter port.

* Connect Arduino USB to PC. Arduino powers up, LED is red.

* Connect 9VDC power to Arduino. LED is green.

* Launch listening application, like TeraTerm or PrintCapture. Ensure your application

*  is set to the Arduino COM port 6, 115200, 8, none, 1.

* Slide switch to ON. (data stream commences)

* Immediately press PLOT key on scope.

* When scope plot completes immediately slide switch to OFF. (data stream halts)

* Reset Arduino in order to start the next plot.

****************************************************************************************|

*/

//#define __serial_plotter // uncomment to calibrate with Arduino Serial Plotter tool

//#define __dummy_data // uncomment to use dummy data for testing

// The pins

#define TestPin 8      // Digital pin 8

#define SwitchPin 9    // Digital pin 9

#define LEDGreenPin 10 // Digital pin 10

#define LEDRedPin 11   // Digital pin 11

#define PenInputPin 12 // Digital pin 12

#define IdleLED 13     // Digital pin 13

#define X_voltage 0    // Analog pin A0

#define Y_voltage 1    // Analog pin A1

#define VinVoltage 2   // Analog pin A2

// The numbers

unsigned long LoopTime; // Marker for loop() delay timer

unsigned int PenState = HIGH;        // Default pen up

unsigned int OldPenState = PenState; //

const unsigned int numReadings = 8;       // Must be a power of two so we can right shift later

unsigned int Xreadings[numReadings];      // Arrays for computing moving average for smoothing

unsigned int Yreadings[numReadings];      //

unsigned int Xtotal = numReadings * 511;  // Arrays will be preloaded

unsigned int Ytotal = Xtotal;             //

unsigned int readIndex = 0;               //

const unsigned int dcXOffset = 0; //21;  // Compensate for any DC offset. Adjust this to get approximately 511,511.

const unsigned int dcYOffset = 0; //23;  //

// The strings

String ADCXStr; // ADC X counts string

String ADCYStr; // ADC Y counts string

String CmdStr;  // HPGL command string

#ifdef __dummy_data

const float slew = 614.0; // max slew rate is 3.0 volts/second = 614 counts/second (@5 volts = 1023 counts)

const float Freq = (1.0/PI) * (slew / 1023.0);  // This frequency sinusoidal will have that slew rate (at the zero crossing)

const float TWO_PI_F = 2.0 * PI * Freq;

#endif

/*

* The Setup ****************************************************************************|

*/

void setup() {

// Set I/O pins

pinMode(TestPin, OUTPUT); // timing test pin

pinMode(IdleLED, OUTPUT);

pinMode(LEDRedPin, OUTPUT);

pinMode(LEDGreenPin, OUTPUT);

pinMode(SwitchPin, INPUT_PULLUP);

pinMode(PenInputPin, INPUT_PULLUP); // Pen signal is from Normally Open (N.O.) relay contacts

analogReference(DEFAULT);

// Set LEDs

digitalWrite(IdleLED, HIGH); // Turn on 'idle' LED

digitalWrite(LEDRedPin, LOW); // Turn red/green LED off

digitalWrite(LEDGreenPin, LOW);

// initialize serial communication:

Serial.begin(115200);

// Check for 9V Vin external power, stay in setup() until 9V supply connected

/*

* Why 900 counts? To gaurantee that a 7.5 VDC±5% power pack will not trigger the

* 'Go' signal under maximal conditions:

* 7.5v+5%; R1Ω-1%; R2Ω+1%; USB 5v-5% (AREF); +3 counts (2 LSB ADC error)

*/

while (analogRead(VinVoltage) < 900) {

digitalWrite(LEDRedPin, HIGH); // Turn LED red (no-go)

digitalWrite(LEDGreenPin, LOW);

delay(100);

} // Wait here until user plugs in a 9VDC, or greater, power supply

// Vin OK, continue

digitalWrite(LEDRedPin, LOW); // Turn LED green (go)

digitalWrite(LEDGreenPin, HIGH);

//Wait here for switch to be switched on (ground the pin)

while (digitalRead(SwitchPin) == HIGH) {

delay(100);

}

// User is ready for data stream, let 'er rip!

analogReference(EXTERNAL); // External precision 5.00V reference

// Preset the averaging arrays

for (int thisReading = 0; thisReading < numReadings; thisReading++) {

Xreadings[thisReading] = 511;

Yreadings[thisReading] = 511;

}

digitalWrite(IdleLED, LOW); // turn off 'idle' LED, transmission will commence (TX LED will light)

// HP-GL plot header commands

Serial.println("IN;"); // Initialize

Serial.println("PS8900,8900;"); // Plot size; A-size

Serial.println("IP0,0,8900,8900;"); // Input Scaling Points P1 & P2; A-size drawing (8.5" x 11", less margins)

Serial.println("SC0,1023,0,1023,1;"); // Scale; Isotropic

//Serial.println("OH;"); // Output hard clip limits

//Serial.println("OP;"); // Output P1 & P2

Serial.println("PU;"); // Pun Up

Serial.println("SP1;"); // Select Pen 1

} // end setup

/*

* The Loop *****************************************************************************|

*/

void loop() {

LoopTime = micros(); // mark loop start

// Read X, Y and Pen signals as close together as possible and compute a moving average for smoothing

Xtotal -= Xreadings[readIndex]; // Subtract an old value from the total

Ytotal -= Yreadings[readIndex]; //

PenState = digitalRead(PenInputPin); // Read pen input; returns HIGH or LOW

Xreadings[readIndex] = analogRead(X_voltage) - dcXOffset; // compensate for dc offset

delayMicroseconds(16); // Delay two ADC clock cycles (there is less 'jitter' on the Y readings than for a one ADC clock cycle dalay, don't ask me why)

Yreadings[readIndex] = analogRead(Y_voltage) - dcYOffset;

Xtotal += Xreadings[readIndex];   // Add the new value to the total

Ytotal += Yreadings[readIndex++]; // Increment readIndex here

readIndex %= numReadings;  // Wrap around readIndex here

ADCXStr = Xtotal >> 3; // Total divided by numReadings (and convert to string)

ADCYStr = Ytotal >> 3;

//

// If pen has changed state then output pen command

// Pen UP signal is HIGH (relay open, interanl pull up resistor), pen DOWN is LOW (relay closed to ground)

if (PenState != OldPenState) {

(PenState == HIGH) ? (CmdStr = "PU;") : (CmdStr = "PD;");

OldPenState = PenState; // Remember pen state

Serial.println(CmdStr);

}

#ifdef __dummy_data

// dummy data for now, plot a circle

float t = (float)millis() / 1000.0;

float arg = TWO_PI_F * t;

ADCXStr = int(511.5*cos(arg)+511.5);

ADCYStr = int(511.5*sin(arg)+511.5);

#endif

#ifdef __serial_plotter

int Pen;

(PenState == HIGH) ? Pen = 10 : Pen = 0;  // View the Pen up/down signal

CmdStr = "1023 0 " + ADCXStr + " " + ADCYStr + " " + Pen; // 1023 0 prevents autoscaling

#else

// 'Plot Absolute' Ex: PA1023,250;

CmdStr = "PA" + ADCXStr + "," + ADCYStr + ";";

#endif

Serial.println(CmdStr);

// If switch went HIGH, output final commands, stop data stream, the plot is completed, wait here until RESET

if (digitalRead(SwitchPin) == HIGH) {

Serial.println("PU;");

Serial.println("SP;");

Serial.println("IN;");

while (true) {  // wait here for RESET

// flash 'idle' LED to indicate need for reset

digitalWrite(IdleLED, HIGH);

delay(200);

digitalWrite(IdleLED, LOW);

delay(300);

}

}

//

/* Wait here for remainder of loop time.

* Full scale plot swing is 1023 counts in 1.667 seconds (5 volt full swing / 3.0 volts/second).

* So, no more than one plot command every 1.630ms is necessary (1.667/1023). ~614 samples/second.

*/

while (micros() - LoopTime < 1630) {;} // wait here for remainder of loop time

// Short (0.250Âµs) output pulse to test loop timing.

PORTB |= B1; PORTB ^= B1; // digitalWrite(TestPin, HIGH); digitalWrite(TestPin, LOW);

}  // end loop

Part List:

Cost:

| Arduino UNO | $25 |
| --- | --- |
| Arduino Proto Shield (needed for the headers only. I spun my own board) | $8 |
| Parts (25) | $57 |
| Board by Aiser (qty 3) | $34 (€31) |
| Total | $124 (€114) |

References:

•	Atmel ATmega328/P 8-bit AVR Microcontroller Datasheet.

•	Arduino Programming Nootebook by Brian W. Evans.

•	Arduino Uno Rev3 Schematic

•	Tektronix 2232 Digital Storage Oscilloscope User Manual, 070-7066-01

•	Tektronix 2232 Digital Storage Oscilloscope Service Manual, 070-7067-01

•	The HP-GL/2 and HP RTL Reference Guide, A Handbook for Program Developers, 5961-3526